Understanding Polarimetry

Instructor notes

1.	If using computers, the Polarimeter requires Logger Pro version 3.8.4.2 or newer. LabQuest users require LabQuest App version 1.6 or newer. An update can be downloaded from the Vernier web site.

2.	If you are using a LabPro interface, turn off Live Readouts under the Experiment menu prior to starting data collection. This must be done each time a new session is started. Note: This requirement should go away in Logger Pro 3.8.5, available in April 2012.

3.	The Vernier Polarimeter has two cables. One is a BTA connector that should be plugged into CH1 of your interface and one is a BTD connector that should be plugged into DIG1 of your interface. Both connections are required for data collection.

4.	After connecting the Polarimeter to a Vernier interface, open the data-collection software. If using Logger Pro, an experiment file will appear after a few seconds that has a graph of Illumination on the y-axis and Angle on the x-axis. If this does not occur, choose New from the File menu.

5.	Sucrose is often used as a calibration standard for polarimeters, which is why it is suggested in this experiment. You may also use other sugars and should obtain similar results. Literature values for specific rotation were obtained from the CRC Handbook, determined in water at concentrations of 1–5 g per 100 mL of solution, at 20–25°C, and at 589 nm.

6.	There are several ways to locate the angle at which the maximum illumination occurs. This experiment suggests the student use a Gaussian fit first; then, as an extension, to use the other two and compare all three.

Statistics: To simply get the angle with the highest illumination, highlight the peak of interest in Logger Pro or LabQuest App, as shown in Figure 1. Choose Statistics from the Analyze menu. Record the angle value where the illumination is at a maximum, as presented in the box. This method is the fastest and will result in reproducibility of the angle of rotation measurement of ±2.0°.

Gaussian: To improve your accuracy with a better fit, highlight the peak of interest using Logger Pro or LabQuest App, as shown in Figure 1. Then, choose Curve Fit from the Analyze menu. From the list of available Equations, select Gaussian. Select Try Fit in Logger Pro; in LabQuest App, the fit will run automatically. The B coefficient presented represents the angle at maximum illumination. This method will result in reproducibility of the angle of rotation measurement of ±0.3°. The data are not a true Gaussian, but the ease and accuracy of this methodology make it a good option. For best results, be consistent in the way you select your peaks.

Cosine Squared: To incorporate all of your data into the fit, you can fit the data to its true waveform, a cosine squared, in either Logger Pro or LabQuest App. Choose Curve Fit from the Analyze menu. From the list of available General Equations, select Cosine Squared. Select Try Fit in Logger Pro; in LabQuest App, the fit will run automatically. In this fit, the x-value corresponding to the maximum y-value is obtained from the negative of the phase shift parameter, –C. This is a nonlinear fit which undergoes numerous iterations and has the possibility of no convergence, which will result in an unreasonable answer. With all nonlinear fits, it is important to make sure the resulting value is reasonable based on the data presented in the graph. This method is the most time consuming; however, it will result in reproducibility of the angle of rotation measurement of ±0.1°.

7.	If a LabPro interface is being used, allow a few seconds at both the beginning and ending of the 15 second data-collection period without moving the analyzer. Typically, a warning “Waiting for data…” will appear when data collection has been initiated; wait until this message has disappeared before rotating the analyzer. It may be helpful to extend the data-collection time to ensure students do not continue to rotate the analyzer after data collection has ended.

LabPro users have these extended instructions because the optical encoder that tells Logger Pro what angle the analyzer is sitting on does not continuously collect data on the LabPro like it does on the LabQuest and LabQuest Mini interfaces.

8.	Due to the nature of light transmission, it is important that the sample be transparent and homogeneous. The sample can have a moderate amount of color but make sure it is not too dark to ensure that light will still pass through it onto the detector.

9.	The illumination value should not be used as a quantitative value for this device. If you choose to use it for this purpose, note that this value is very sensitive to change. These slight variations can be brought on by several different variables in the sample including inhomogeneity (such as particulates and bubbles), height of the sample in the cell, concentration of the sample, and aberrations in the glass cell or path length.

HAZARD ALERTS

Sucrose: May be harmful if inhaled, absorbed through skin or swallowed. May cause eye irritation. NFPA Rating: Health hazard(0, Fire(0, Reactivity hazard(0.

Hazard information reference: 
Sigma-Aldrich Co., 1-800-325-3010, www.sigmaaldrich.com/safety-center/msds-search.html

COMPOUND INFORMATION

| Compound | Chemical formula | Molar mass (g/moL) | Specific Rotation (°) | Melting temperature (°C) |
| --- | --- | --- | --- | --- |
| sucrose | C12H22O11 | 342.3 | +66.5 | 186 |

Sample DATA

Part I  Exploring Path Length

Angleblank (°) = 197.2

|  | Run 1 | Run 2 | Run 3 | Run 4 | Run 5 |
| --- | --- | --- | --- | --- | --- |
| Sample height (cm) | 2.0 | 4.0 | 6.0 | 8.0 | 10.0 |
| Anglesample (°) | 198.0 | 199.2 | 200.5 | 201.4 | 202.8 |
| Angle of rotation, α (°) = Anglesample – Angleblank | 0.8 | 2.0 | 3.3 | 4.2 | 5.6 |

Part II  Exploring Concentration

Angleblank (°) = 37.9

|  | 10% sample | 20% sample | 30% sample |
| --- | --- | --- | --- |
| Sample height (cm) | 10.1 | 10.0 | 10.1 |
| Anglesample (°) | 43.3 | 50.9 | 57.3 |
| Angle of rotation, α (°) = Anglesample – Angleblank | 5.5 | 13.0 | 19.4 |
| Exact concentration (g/mL) | 0.083 | 0.197 | 0.295 |

answers to data analysis questions

Part I  Exploring Path Length

1.	See graph below.

2.	Linear.

3.	See graph below.

y = mx+b

y = 0.5443x – 0.0

x = 10.0 cm

y = 5.4°

4.

Part II  Exploring Concentration

5.	See table above. Sample calculation:

6.	See graph below.

7.	Linear.

Extension

By simply curve fitting the same set of data in Logger Pro, the cosine squared fit should be the most accurate, if done properly. This method will result in an angle of rotation accuracy of ±0.1°. Fitting with a Gaussian will result in an angle of rotation accuracy of ±0.3°. Statistics will result in an angle of rotation accuracy of ±2.0°.

Sample Graphs

Part I  Exploring Path Length

Optical rotation data for sucrose at varying path lengths

Relationship between sample cell height and corrected angle of rotation for sucrose

Part II  Exploring Concentration

Optical rotation data for 10 cm height sucrose at varying concentrations

Relationship between concentration and angle of rotation for sucrose

# Text boxes

Figure 1  Selection for fits