Determining the Concentration
of a Solution: Beer’s Law

The primary objective of this experiment is to determine the concentration of an unknown cobalt (II) chloride, CoCl2, solution. You will use a Vernier SpectroVis Plus spectrophotometer to measure the concentration of each solution. You will first measure the absorbance of a standard solution over the visible light spectrum and select the wavelength of maximum absorbance (λ max). A higher concentration of the solution absorbs more light (and transmits less) than a solution of lower concentration.

You will prepare five cobalt (II) chloride solutions of known concentration (standard solutions) and measure the absorbance of the standard solutions at the λ max. The graph of absorbance vs. concentration for the standard solutions will describe a direct relationship, known as Beer’s law.

Finally, you will determine the concentration of an unknown CoCl2 solution by measuring its absorbance with a spectrophotometer and using the best-fit line equation of the Beer’s law curve to calculate the unknown’s concentration.

OBJECTIVES

In this experiment, you will

- Measure and analyze the visible light absorbance spectrum of a standard cobalt (II) chloride solution to determine the maximum wavelength of absorbance.

- Prepare and test the absorbance of five standard cobalt (II) chloride solutions.

- Calculate a standard curve from the test results of the standard solutions.

- Test the absorbance of a cobalt (II) chloride solution of unknown molar concentration.

MATERIALS

| Vernier SpectroVis Plus | 0.20 M cobalt (II) chloride, CoCl2, solution |
| --- | --- |
| LabQuest or computer | cobalt (II) chloride, CoCl2, unknown solution |
| LabQuest App or Logger Pro | distilled water |
| pipet pump or pipet bulb | test tube rack |
| one cuvette | stirring rod |
| two 10 mL pipets or graduated cylinders | tissues (preferably lint-free) |
| five 20 × 150 mm test tubes | two 100 mL beakers |

PROCEDURE

1.	Obtain and wear goggles.

2.	Obtain ~40 mL of 0.20 M CoCl2 solution and distilled water in separate beakers.

3.	Label five clean, dry, test tubes 1–5. Use pipets to prepare four standard solutions according to the chart below (the fifth standard is the stock 0.10 M CoCl2 solution). Thoroughly mix each solution with a stirring rod. Clean and dry the stirring rod between uses.

| Test Tube number | 0.20 M CoCl2 (mL) | Distilled H2O (mL) | Concentration (M) |
| --- | --- | --- | --- |
| 1 | 10 | 0 | 0.20 |
| 2 | 8 | 2 | 0.16 |
| 3 | 6 | 4 | 0.12 |
| 4 | 4 | 6 | 0.08 |
| 5 | 2 | 8 | 0.04 |

4.	Connect the SpectroVis Plus spectrophotometer to the USB port of LabQuest or a computer. Start the data-collection program, and then choose New from the File menu.

5.	Calibrate the spectrophotometer.

Prepare a blank by filling an empty cuvette ¾ full with distilled water.

Choose Calibrate from the Sensors menu of LabQuest or the Experiment menu of Logger Pro.

Place the blank in the spectrophotometer; make sure to align the cuvette so that the clear sides are facing the light source of the spectrometer. When the warmup period is complete, select Finish Calibration. Select OK.

6.	Determine the maximum wavelength for CoCl2 (aq) and set up the data collection mode.

- Empty the blank cuvette and rinse it twice with small amounts of the 0.20 M CoCl2 solution in Test Tube 1. Fill the cuvette ¾ full with the 0.20 M CoCl2 solution and place it in the spectrophotometer.

Start data collection. A full spectrum graph of the CoCl2 solution will be displayed. Note that one area of the graph contains a peak absorbance. Stop data collection.

Store the run by tapping the File Cabinet icon in LabQuest, or choosing Store Latest Run from the Experiment menu in Logger Pro.

To set up the data collection mode:

- In LabQuest, the wavelength of maximum absorbance (( max) is automatically identified. Tap the Meter tab, then tap Mode. Change the mode to Events with Entry. Enter the Name (Concentration) and Units (mol/L). Select OK.

- In Logger Pro, click the Configure Spectrometer button, . Click Abs vs. Concentration as the Collection Mode. The wavelength of maximum absorbance (( max) will be selected.

7.	Collect absorbance-concentration data for the five standard solutions.

- Leave the cuvette in the spectrophotometer and start data collection. When the absorbance reading stabilizes, select Keep. Enter “0.20” as the concentration of the solution and select OK. The absorbance and concentration values have now been saved for the first solution.

- Discard the cuvette contents as directed. Using the solution in Test Tube 2, rinse and fill the cuvette ¾ full. Wipe the cuvette and place it in the spectrometer. When the absorbance reading stabilizes, select Keep. Enter “0.16” as the concentration.

Repeat Step 7b for the remaining test tubes of the standard CoCl2 solution. When you have finished testing the standard solutions, stop data collection.

8.	To determine the best-fit line equation for the CoCl2 standard solutions, choose Curve Fit from the Analyze menu. Choose Linear as the Fit Equation. Write down the equation for the standard solutions in your data table.

9.	Determine the concentration of the unknown CoCl2 solution.

- Obtain about 5 mL of the unknown CoCl2 solution.

- Rinse the cuvette twice with the unknown solution and fill it about ¾ full. Wipe the outside of the cuvette and place it into the spectrophotometer.

- Monitor the absorbance value on the displayed Meter in Logger Pro (or tap the Meter tab in LabQuest). Write down the absorbance of the unknown in your data table.

- Choose Interpolate from the Analyze menu. To determine the concentration of the unknown, interpolate along the regression line to convert the absorbance value of the unknown to concentration.

Write down the absorbance and the concentration of the unknown in your data table.

10.	Dispose of any of the remaining solutions as directed.

DATA TABLE

| Test Tube | Concentration (mol/L) | Absorbance |
| --- | --- | --- |
| 1 | 0.02 |  |
| 2 | 0.04 |  |
| 3 | 0.06 |  |
| 4 | 0.08 |  |
| 5 | 0.10 |  |
| 6 | Unknown number ____ |  |

DATA ANALYSIS

Your lab report should contain the following information:

- a graph showing the data and linear-regression equation for the standard solutions

- the molar concentration of your unknown CoCl2 solution

- an explanation of your calculation of the unknown’s concentration