Nucleic Acid Quantitation

Quantitation of nucleic acids is commonly performed to determine the concentration of DNA or RNA present in a mixture. It is also used to determine the purity of a DNA or RNA sample. Contamination of a nucleic acid solution by proteins, carbohydrates, and other organic molecules can be determined using a procedure called the A260/280 ratio.

The basis of this test relies on the Beer-Lambert Law: A = εbc; where A is absorbance, ε is the molar extinction coefficient, b is the cell path length, and c is the sample concentration. The commonly accepted extinction coefficients for a 1 mg/mL nucleic acid solution at 260 nm and 280 nm are 20 and 10, respectively. In proteins, the extinction coefficient values at 260 nm and 280 nm at a concentration of 1 mg/mL are 0.57 and 1.00, respectively. Therefore, nucleic acid samples would be expected to have a higher absorbance at 260 nm than at 280 nm; in a protein sample, the opposite is true. Using these extinction coefficients, pure nucleic acid samples would have an A260/280 ratio of 2.0, while pure protein samples would have an A260/280 ratio of 0.57.

OBJECTIVES

In this experiment, you will determine the A260/280 ratio of a sample.

MATERIALS

| LabQuest or computer | prepared biological sample containing |
| --- | --- |
| LabQuest App or Logger Pro | DNA, RNA and/or protein |
| Vernier UV-VIS Spectrophotometer | buffer solution used to prepare |
| quartz cuvette(s) | biological sample |

PROCEDURE

1.	Obtain and wear goggles and gloves.

2.	Connect the AC power supply to the Vernier UV-VIS Spectrophotometer. Turn the power switch to the ON position.

3.	When the LED for the lamp indicator stays green, connect the spectrophotometer to the USB port of LabQuest or a computer. Start the data-collection program, and then choose New from the File menu.

4.	Calibrate the spectrophotometer.

- To calibrate the spectrophotometer, choose Calibrate ► Spectrometer from the Sensors menu of LabQuest or the Experiment menu of Logger Pro. Note: For best results, allow the spectrophotometer to warm up for a minimum of 10 minutes.

- Prepare a blank by filling the empty quartz cuvette 3/4 full with the buffer solution used to prepare your sample.

- After the spectrophotometer has warmed up, place the cuvette in the spectrophotometer. Align the cuvette so that one of its clear sides is facing the light source.

- Select Finish Calibration. Select OK.

5.	Collect absorbance vs. wavelength data.

- Fill a cuvette with ~3 mL of the biological sample to be tested. Your sample should not be turbid or have any particulates in it.

- Place the sample in the spectrophotometer.

- Start data collection. Once the spectrum is displayed, stop data collection.

- Make sure the absorbance values at 260 nm and 280 nm are between 0.1 and 1.0 absorbance units. Any values outside this range may introduce an error. If your sample is outside this range, dilute or concentrate it and take another spectrum.

6.	Using the Examine feature in Logger Pro or LabQuest App, or by scrolling down the data table, find the exact absorbance values at 260 nm, 280 nm, and 340 nm. Record the absorbance values in your data table. Note: A value of 340 nm is used to normalize here because there can be contaminants in your solution that scatter the light and create an absorbance offset throughout this wavelength region.

7.	Store the run. In Logger Pro, do this by choosing Store Latest Run from the Experiment menu. In LabQuest App, you can store a run by tapping the file cabinet icon.

8.	Repeat Steps 5–7 two more times (optional).

DATA table

| Run | A260 | A280 | A340 | A260/280 (corrected) |
| --- | --- | --- | --- | --- |
| 1 |  |  |  |  |
| 2 |  |  |  |  |
| 3 |  |  |  |  |
|  |  |  | Average |  |

DATA Analysis

A260/280(corrected)=A260−A340A280−A3401.   Correct your A260 and A280 absorbance values for the A340 absorbance value according to the following equation:

Complete this calculation for each run. Record these values in the data table.

2.   Calculate the average A260/280 of your three trials. Record this value in the data table.

3.	Discuss the purity and composition of your samples.